## Slide 1: Setup a new measuring instance

## Slide 2: Bill of Material
Pi 5 with at least 4GB RAM
NVME Extension and minimum 128GB SSD
abaton clima measurement system PCB [short: ACMS]
Network Cable RJ45
Digisense
  Mux PCB
  2 x Scale Duett 8kg
9V PSU
6 x ENS 210
4 x DS18B20 (PT1000)
2 x Hailege MLC90614 IR Cam
Webcam with USB
2 x flow measurement unit

## Slide 3: Setting up the new pi
take a new pi and generate an image for a pi5 with some settings:
  username: rabaton
  password: lack..
  wifi: rudi // bbaton125!
  enable ssh
first time boot use an external monitor and obtain the IP address of the device
  enable under preferences -> raspberry configuration
  onewire
  I2C
  SPI
then reboot. afterwards -> Accessoires -> SD Card Copier
  copy content of SD card onto SSD (new UUID Partitions must be checked)
  open sudo nano /boot/firmware/config.txt and add the lines:
    dtparam=pciex1
    dtoverlay=w1.gpio -> dtoverlay=w1.gpio,gpiopin=26
power off, remove SD card, power on
register the pi with our raspberry cloud:
  sudo apt update
  sudo apt upgrade
  click icon right top on the desktop with the two squares and the circles and use
    beschaffung@abaton.studio
    rasART135!
ssh key für ds_readout git raphael geben
  ssh-keygen -t rsa
  den public key ins gitHub kopieren als deploy key
  git clone https://github.com/xtlc/ds_readout.git machen am pi
  cd ds_readout
  python -m venv venv
  source venv/bin/activate
  python -m pip install -r req
  Create an empty env file nano .env, save with CRTL + O and [y]
  echo 'export GIT_MERGE_AUTOEDIT=no' >> $HOME/.bash_profile && . $HOME/.bash_profile #make sure the git pull does not open an editor

## Slide 4: Setting up telegraf
Telegraf is a service that makes sure, when data should be stored at influx but the internet connection is disrupted, that the data is stored locally and sent once internet is available again.
Install guide: Link and quick start: Link
SSH into the raspberry or access locally:
  wget -q https://repos.influxdata.com/influxdata-archive_compat.key
  echo '393e8779c89ac8d958f81f942f9ad7fb82a25e133faddaf92e15b16e6ac9ce4c influxdata-archive_compat.key' | sha256sum -c cat influxdata-archive_compat.key | gpg --dearmor | sudo tee /etc/apt/trusted.gpg.d/influxdata-archive_compat.gpg > /dev/null
  echo 'deb [signed-by=/etc/apt/trusted.gpg.d/influxdata-archive_compat.gpg] https://repos.influxdata.com/debian stable main' | sudo tee /etc/apt/sources.list.d/influxdata.list
  sudo apt-get update && sudo apt-get install telegraf
  sudo nano /etc/telegraf/tg.conf (be sure to input influx token)
Start telegraf with raspberry (re)boot
  sudo nano /etc/systemd/system/telegraf.service -> copy config from telegraf service file
  sudo systemctl daemon-reload (you need to execute this if you made a typo somewhere in the confs to reload them)
  sudo systemctl start telegraf
    Check if it really runs: sudo journalctl –u telegraf.service
  Reboot and check if it is running: sudo journalctl –u telegraf.service

## Slide 5: Setting up the scales
Connect MUX to ACMS using a RJ45 cable
Plug 9V power supply into ACMS
Take Sticker from MUX plate and use number
write number into the .env file:
  MUX=0020241017114723 (in this case)
  Connect the left scale to port 1 of the mux
  Connect the right scale to port 2 of the mux
Test the scales with activated venv:
  python scales.py -> you should see something like: 1.903successful, revision of the mux: #21gr002024101711472334

## Slide 6: Installing the mlx90640 IR cam
Cut power to ACMS board and pi before proceeding!
Connect SDA, SCL, VCC, GND as written on the cam and ACMS